# Annex 1

# Validation process of the Real Time PCR for the identification of Phyllosticta citricarpa (van Gent-Pelzer et al., 2007)

Diagnostic protocol

Samples and DNA extraction

The validation experiments were carried out on a series of target referenced samples, consisting of 4 different strains of Phyllosticta citricarpa and 3 non-target samples (table 1). DNA was extracted from fungal cultures and lemon fruit using the commercial plant/fungi DNA extraction kit NucleoSpin Plant II (Macherey-Nagel), according to the manufacturer’s instructions. The quantification was performed by the fluorometer Qubit .1 (Invitrogen). The samples were prepared adding a known amount of target and no-target DNA to the DNA extracted from fruit lemon. The quality and amplification ability of each DNA sample was estimated through a PCR using the universal primer ITS4 and ITS5 (White et al., 1990).

Table 1. Target and non-target fungal strains used for the validation

| Identification name | Reference | Species | Host |
| --- | --- | --- | --- |
| ER1758 | *CBS 828.97 | Phyllosticta citricarpa | Citrus aurantium |
| ER1759 | CBS 102374 | Phyllosticta citricarpa | Citrus aurantium |
| ER1760 | CBS 127455 | Phyllosticta citricarpa | Citrus sinensis |
| ER1761 | CBS 127451 | Phyllosticta citricarpa | Citrus reticulata |
| ER1762 | CBS 120490 | Phyllosticta capitalensis | Citrus paradisi |
| ER1800 | **ITS4xITS5 sequencing | Mycosphaerella sp. | Citrus limon |
| ER1891 | ITS4xITS5 sequencing | Phyllosticta citriasiana | Citrus maxima |

* from CBS, Centraalbureau voor Schimmelcultures-Fungal Biodiversiry Centre (Utrecht, The Netherlands).

**White et al., 1990

PCR amplification

PCR reactions were performed using the forward primer GcF1 (5’- GGT GAT GGA AGG GAG GCC T -3’), the reverse primer GcR1 (5’- GCA ACA TGG TAG ATA CAC AAG GGT -3’) and the dual labeled hydrolysis probe GcP1 (5’- AAA AAG CCG CCC GAC CTA CCT TCA -3’) FAM- and TAMRA-labeled at 5’and 3’ end respectively. Two l of DNA template per reaction were amplified.

The volume and concentrations of each reagent are shown in table 2.

Table 2 Volumes and concentrations of reagents in Real Time PCR

| Reagent | Working concentration | Volume per reaction (µL) | Final concentration |
| --- | --- | --- | --- |
| Molecular grade water |  | 10.2 |  |
| Citrus limon DNA | 30 pg/µL | 1 | 1 pg/µL |
| Master MIX | 2X | 15 | 1X |
| Forward primer GcF1 | 10 µM | 0.75 | 025 µM |
| Reverse primer GcR1 | 10 µM | 0.75 | 0.25 µM |
| Taqman probe GcP1 | 10 µM | 0.3 | 0.1 U/µL |
| DNA |  | 2 |  |
| Total |  | 30 |  |

Thermocycler conditions:

|  | Temperature | Time | Number of cycles |
| --- | --- | --- | --- |
| Initial Denaturation | 94 °C | 10’ | 1 |
| Denaturation | 94 °C | 15” | 40 |
| Annealing/ Elongation | 65 °C | 30” |  |

The fluorescence reading occurred after each annealing/elongation step.

All reagents, materials and instruments used for the validation procedures are reported in table 3

Table 3 Reagents, materials and instruments used for the validation

| Reagents/Material/Instruments | Company |
| --- | --- |
| DNA polymerase: Maxima Probe/ROX qPCR Master Mix (2x) | Thermo Scientific |
| Real time PCR plates: Multiplate 96-Well Unskirted PCR Plates; Microseal ‘B’ Seal Seals | Bio-Rad |
| Filter tips: Standard |  |
| Microfuge tubes: Cliklok Microcentrifuge Tubes | Simport |
| Pipettes: PIPETMAN P10, P20, P100, P200, and P1000 | Gilson |
| Thermocycler: CFX-96 | Bio-Rad |

RING TEST:

The robustness of this method was verified  through a “Ring Test” among 6 laboratories. Six positive samples (3 containing the target DNA slightly above the relative limit of detection and 3 containing the target DNA ten times the relative limit of detection) and 6 negative samples (3 containing no DNA and 3 containing DNA of non-target strains ten times the relative limit of detection) were tested. The results showed:

-100% relative sensitivity

-100% relative specificity

-100% repeatability

-100% reproducibility

Reagents and instruments used for the Ring Test in each laboratories are reported in table 4.

|  | DNA polymerase | Real Time Thermocycler | Software for data analysis |
| --- | --- | --- | --- |
| LAB01 | 2x SuperMix (Bio-Rad) | Bio-Rad CFX | Bio-Rad CFX manager 3.0 |
| LAB02 | TaqMan 2X Universal PCR Master Mix (Applied Biosystem) | Step One Plus Applied Biosystems | Step One Software, for StepOne and StepOnePlus; v2.2.2 |
| LAB03 | TaqMan 2X Universal PCR Master Mix (Applied Biosystem) | 7300 Real Time PCR System, AppliedBiosystem | 7300 SDS Software, v1.3 |
| LAB04 | 2X sSO Fast SuperMix, Biorad | Bio-Rad CFX | Bio-Rad CFX manager 3.0 |
| LAB05 | Platinum qPCR super Mix-UDG (Invitrogen, 11730-025) | Light Cycler 480, Roche | Light Cycler 480 Software release 1.5.0 SP3 (1.5.0.39) |
| LAB06 | Super Mix 2x, Bio-Rad | 7500 Real Time PCR System, AppliedBiosystem | 7500 Software, v2.0.4) |

Table 4 Reagents and instruments used for the Ring Test in each laboratories.

Validation process

As mentioned in ISO 17025 in the laboratory have been recorded the results obtained and the procedure used for the validation. In each amplification experiments a negative isolation control (NIC) was included, to monitor cross-reactions with the host tissue and/or contamination during nucleic acid extraction: nucleic acid extraction and subsequent amplification of a sample of uninfected host tissue or clean extraction buffer.

Analytical sensitivity:  performed on 3 different target strains (ER1758, ER1759, ER1760) with 8 1:10 serial dilutions (from 10 ng to 1x10-6 ng).

Analytical specificity:  performed on 3 different target strains (ER1758, ER1759, ER1760) and 3 different non-target strains (ER1762, ER1800, ER1891) using 1x10-3 ng of DNA template.

Repeatability:  performed on the same isolate in trhee identical experiments at above (1x10-3 ng of DNA template) and at the limit (1x10-5 ng) of detection, with 8 replicates.

Reproducibility:  performed on the same isolate in trhee different experiments at above (1x10-3 ng of DNA template) and at the limit (1x10-5 ng) of detection, with 8 replicates.